MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 60

| Priority | 3 | Floor | 1F-OT | Quantity | 1EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Ultrasonic Surgical Aspirator |  | Installation Location | Operating Room 5 |
| Specification Level |  | High |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The ultrasonic surgical aspirator system shall be used for selective fragmentation, emulsification, irrigation, and aspiration of soft tissues during surgical procedures including neurosurgery, hepatobiliary surgery, and general surgical tumor resection procedures. - The system shall support precise tissue removal with selective soft-tissue fragmentation designed to preserve surrounding vessels, ducts, and critical anatomical structures where clinically applicable. - The equipment shall be suitable for continuous operating room use and compatible with standard operating room suction and irrigation systems and infrastructure. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Microprocessor-controlled ultrasonic surgical console with integrated irrigation and suction control - Stable and continuous output suitable for prolonged procedures - Medical-grade construction resistant to hospital disinfectants - Easy-to-clean surfaces - Compact footprint suitable for OR space - Mobile trolley or cart-mounted configuration - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Ultrasonic Fragmentation Performance [MUST] |  |  |  |  |  |
| - Operating ultrasonic frequency approximately within the 20–40 kHz range or manufacturer-optimized equivalent suitable for ultrasonic tissue fragmentation - Adjustable power/amplitude with multiple selectable levels (minimum 5 steps) - Stable amplitude under load without significant drop in output - Selective soft-tissue fragmentation capability designed to preserve surrounding critical structures where clinically applicable - Smooth and consistent tip oscillation without excessive vibration - Continuous operation capability for minimum 60 minutes without overheating |  |  |  |  |  |
| - Irrigation System Performance [MUST] |  |  |  |  |  |
| - Adjustable irrigation flow rate suitable for continuous ultrasonic surgical irrigation (approximately up to 500–800 mL/min or equivalent) - Stable and continuous fluid delivery - Independent irrigation control - Compatible with standard sterile saline bags - Easy tubing replacement during surgery |  |  |  |  |  |
| - Aspiration/Suction Performance [MUST] |  |  |  |  |  |
| - Adjustable suction vacuum up to approximately –600 mmHg - Independent suction control separates from ultrasonic power - Efficient debris removal without clogging - Integrated clog-resistant aspiration channel |  |  |  |  |  |
| - Handpiece & Ergonomics [MUST] |  |  |  |  |  |
| - Lightweight ergonomic handpiece suitable for prolonged use - Handpiece weight ≤ 300 g (approx.) - Low vibration and low noise design - Low-noise operation suitable for operating room environment (manufacturer specification) - Comfortable grip with non-slip surface - Quick-connect mechanism for tips - Cable strain relief to prevent damage |  |  |  |  |  |
| - Tips & Sterilization [MUST] |  |  |  |  |  |
| - Interchangeable surgical tips for various procedures - Minimum 5 assorted tip types included - Autoclavable handpiece and tips - Compatible with steam sterilization at ≥134°C - Designed to withstand ≥ 500 sterilization cycles without degradation - Waterproof/sealed design to prevent fluid ingress |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Automatic over-temperature protection - Tip clog detection or overload alarm - Fluid leakage protection - Audible and visual alarm indicators - Automatic shutdown in abnormal conditions - Electrical leakage protection compliant with IEC 60601-1 |  |  |  |  |  |
| - Display & Control Interface [MUST] |  |  |  |  |  |
| - Digital display showing power level, suction level, and irrigation settings - Easy-to-use control panel - Footswitch activation with variable control for ultrasonic output - Clear status indicators |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate surgical use: - Main console unit – 1 - Mobile trolley/cart – 1 - Ultrasonic handpiece – 1 - Footswitch – 1 - Surgical tips – minimum 5 assorted - Irrigation tubing sets – minimum 10 - Suction tubing sets – minimum 10 - Sterilization tray or case – 1 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and on-site installation - Functional testing of ultrasonic, irrigation, and suction systems - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site clinical training for surgeons and nurses - Technical training for biomedical engineers - Operation, cleaning, sterilization, and troubleshooting included |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel costs - Preventive maintenance included during warranty period - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking and FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/