MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 61

| Priority | 3 | Floor | 2F, 3F | Quantity | 2EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | ECG Machine |  | Installation Location | Treatment Room (2F) Treatment Room (3F) |
| Specification Level |  | Low |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The electrocardiograph system shall be used for acquisition, display, recording, analysis, and printing of standard 12-lead electrocardiograms for clinical evaluation of cardiac rhythm and myocardial abnormalities. - The equipment shall support routine clinical screening, emergency assessment, and inpatient diagnostic assessment for adult and pediatric patients. - The system shall provide accurate signal acquisition with high noise immunity and optional interpretive analysis support. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Medical-grade diagnostic ECG machine (NOT consumer or handheld type) - Simultaneous 12-lead acquisition and recording - Portable tabletop design with integrated display and printer - Lightweight and easy mobility between departments - Durable housing resistant to hospital disinfectants - Easy-to-use interface suitable for nurses and technicians - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Signal Acquisition Performance [MUST] |  |  |  |  |  |
| - Simultaneous acquisition of all 12 leads - Sampling rate ≥ 500 samples/second/channel - Frequency response: 0.05–150 Hz (diagnostic standard) - Input impedance ≥ 10 MΩ - Common Mode Rejection Ratio (CMRR) ≥ 100 dB - Internal noise level compliant with IEC 60601-2-25 diagnostic requirements - Baseline stability suitable for low-frequency drift reduction - Protection against defibrillation and electrosurgical interference |  |  |  |  |  |
| - Measurement & Analysis Functions [MUST] |  |  |  |  |  |
| - Automatic ECG measurement with optional interpretive analysis support - Heart rate, PR interval, QRS duration, QT/QTc, axis calculation - Arrhythmia detection capability - Manual and automatic modes selectable - Pediatric and adult algorithm support |  |  |  |  |  |
| - Recording & Printing Performance [MUST] |  |  |  |  |  |
| - Built-in high-resolution thermal printer - Printer resolution ≥ 200 dpi - Paper width: standard 80–110 mm - Selectable paper speeds: 5, 10, 25, 50 mm/s - Selectable sensitivities: 2.5, 5, 10, 20 mm/mV - Multiple recording formats (3/6/12 channel print) - Real-time preview before printing |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Color LCD display ≥ 7 inches - Simultaneous waveform preview of all leads - Touchscreen or keypad interface - Adjustable brightness - Patient demographic data entry (ID, name, age, gender) - Multi-language support (English mandatory) |  |  |  |  |  |
| - Data Storage & Connectivity [MUST] |  |  |  |  |  |
| - Internal memory storage ≥ 500 ECG records - ECG recall and review capability - USB port for data export (PDF/image/raw format) - LAN connectivity shall be provided; Wi-Fi connectivity acceptable - DICOM 3.0 compatibility or equivalent hospital integration - Date/time stamping of all records |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Internal rechargeable battery - Battery backup ≥ 3 hours continuous operation - Battery level indicator and low battery alarm |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Patient leakage current protection compliant with IEC 60601-1 - Defibrillator protection - Lead-off detection alarm - Automatic self-test at startup - Electrical safety protection against overload |  |  |  |  |  |
| - Mechanical & Portability [MUST] |  |  |  |  |  |
| - Compact and lightweight design (≤ 5 kg preferred) - Integrated carry handle - Stable base or cart option - Durable keyboard/buttons for frequent clinical use |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical operation: - ECG main unit – 1 - 12-lead patient cable – 1 - Limb electrodes (reusable clips) – 4 - Chest electrodes (suction type) – 6 - ECG paper rolls – minimum 20 - Rechargeable battery (installed) - Power cable – 1 - Carry case or protective cover – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site delivery and functional testing - Verification of signal quality and printing accuracy - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for nurses and technicians - Training includes patient setup, recording, interpretation, cleaning, and maintenance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel - Preventive maintenance included during warranty - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-2-25 (Particular requirements for electrocardiographs) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified |  |  |  |  |  |

/End/