MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 62

| Priority | 3 | Floor | 2F, 3F | Quantity | 2EA for each Treatment Room |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Low Vacuum Aspirator |  | Installation Location | Treatment Room (2F) Treatment Room (3F) |
| Specification Level |  | Low |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The low vacuum aspirator shall be used for controlled suction of fluids, secretions, and minor surgical debris during bedside procedures, wound care, minor surgical procedures, and emergency treatment. - The system shall provide stable and adjustable negative pressure with safe and low-noise operation suitable for continuous use in wards and treatment rooms. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Electrically powered medical suction unit designed for hospital clinical use - Continuous-duty operation suitable for prolonged bedside procedures - Portable and compact configuration - Low-noise operation suitable for patient areas - Disinfectant-resistant housing and smooth cleanable surfaces - Easy operation for nurses and clinical staff - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Suction Performance [MUST] |  |  |  |  |  |
| - Adjustable vacuum range suitable for low to moderate clinical suction applications (approximately up to –60 kPa or equivalent) - Flow rate capacity ≥ 20 L/min - Fine and continuous vacuum adjustment via regulator knob - Stable suction performance without fluctuation - Accurate and easy-to-read vacuum gauge displaying kPa and/or mmHg |  |  |  |  |  |
| - Continuous Operation Capability [MUST] |  |  |  |  |  |
| - Designed for continuous operation ≥ 60 minutes without overheating - Stable suction under load - Motor protection against excessive temperature |  |  |  |  |  |
| - Mechanical & Acoustic Performance [MUST] |  |  |  |  |  |
| - Low-noise operation suitable for ward environments (approximately ≤ 65 dB) - Lightweight portable design (≤ 6 kg preferred) - Carry handle or compact structure - Stable base or anti-slip feet - Optional trolley or pole-mount acceptable |  |  |  |  |  |
| - Collection System [MUST] |  |  |  |  |  |
| - Autoclavable or sterilizable collection jars - Minimum capacity per jar: ≥ 1.5 liters - Minimum two (2) collection jars included - Graduated volume markings - Secure lid with leak-proof seal - Easy removal and cleaning |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Overflow protection with float valve to prevent fluid ingress into pump - Hydrophobic bacterial/viral filter - Anti-backflow protection - Automatic motor shut-off in case of blockage or overheating - Electrical leakage protection compliant with IEC 60601-1 - Audible/visual indicator for abnormal condition (if available) |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Internal rechargeable battery backup - Battery operation ≥ 30 minutes for emergency transport or power outage - Battery level indicator |  |  |  |  |  |
| - Controls & Interface [MUST] |  |  |  |  |  |
| - Simple on/off switch - Vacuum adjustment knob - Clear pressure gauge display - Easy-to-clean control panel |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Main suction unit – 1 - Collection jars (≥1.5 L each) – 2 - Silicone suction tubing sets – minimum 2 - Hydrophobic bacterial filters – minimum 10 - Suction catheters/adapters – starter set - Power cable – 1 - Carry handle or trolley (if applicable) – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site functional testing - Verification of suction level and overflow protection - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site user training for nurses and clinical staff - Training includes safe suction operation, cleaning, and maintenance |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel costs - Preventive maintenance included during warranty - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/