MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No. 63

| Priority | 3 | Floor | 2F, 3F | Quantity | 2EA (1 EA for each Treatment Room |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | High-Flow Oxygen System |  | Installation Location | Treatment Room (2F) Treatment Room (3F) |
| Specification Level |  | Mid |  | Remarks |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The high-flow oxygen therapy system shall deliver heated and actively humidified oxygen/air mixtures at high flow rates via nasal cannula interface to provide non-invasive respiratory support for adult and pediatric patients with hypoxemia or respiratory distress. - The system shall provide accurate oxygen concentration, stable flow delivery, and adequate humidification to support patient comfort and airway conditioning during therapy. - The equipment shall be suitable for continuous use in emergency and critical care settings. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Integrated high-flow oxygen therapy system with flow generator, oxygen blending, heating, and humidification - Continuous-duty operation suitable for ICU/ER use - Compact and mobile design suitable for bedside use - Durable housing resistant to hospital disinfectants - Easy-to-use interface for nurses and respiratory therapists - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Flow Performance [MUST] |  |  |  |  |  |
| - Adjustable total flow range: minimum 2–60 L/min or wider - Flow accuracy: ±5% or better - Stable flow delivery without fluctuation during operation - Rapid response to flow setting changes - System shall be capable of delivering high-flow heated and humidified respiratory gas therapy via nasal cannula interface |  |  |  |  |  |
| - Oxygen Concentration (FiO₂) Performance [MUST] |  |  |  |  |  |
| - Adjustable FiO₂ range: 21%–100% - FiO₂ accuracy: ±3% or better - Integrated air/oxygen blender - Stable oxygen concentration independent of flow changes |  |  |  |  |  |
| - Heating & Humidification Performance [MUST] |  |  |  |  |  |
| - Integrated heated humidifier system - Temperature control range: approximately 31–37°C or wider - Temperature accuracy: ±1°C or better - Uniform heated humidification designed to support airway comfort during high-flow therapy - Heated breathing circuit to minimize condensation - Water chamber capacity ≥ 300 mL - Easy refill and cleaning design |  |  |  |  |  |
| - Patient Comfort & Clinical Safety [MUST] |  |  |  |  |  |
| - Stable temperature and humidity to improve patient tolerance - Suitable for prolonged therapy - Compatible with adult and pediatric nasal cannula interfaces - Low resistance breathing circuit |  |  |  |  |  |
| - Display & User Interface [MUST] |  |  |  |  |  |
| - Digital color display - Real-time display of: - Flow rate - FiO₂ - Temperature - Alarm status - Clear control knobs or touchscreen - Simple parameter setting - Visual and audible alarm indicators |  |  |  |  |  |
| - Alarm & Safety Features [MUST] |  |  |  |  |  |
| - Audible and visual alarms for: - Oxygen supply failure - Disconnection or circuit leak - High/low temperature - Low water level - System malfunction - Automatic shutdown or safe mode in abnormal conditions |  |  |  |  |  |
| - Mechanical & Acoustic Performance [MUST] |  |  |  |  |  |
| - Low-noise operation suitable for ICU environment (manufacturer specification) - Stable base or mobile stand with lockable wheels - Lightweight and easy to reposition |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Internal rechargeable battery backup - Battery runtime ≥ 30 minutes to allow patient transport or short power failure - Battery level indicator and alarm |  |  |  |  |  |
| - Connectivity & Data (if available) [MUST/ACCEPTABLE] |  |  |  |  |  |
| - Usage hours display - Basic therapy log or event history - USB or LAN port for service/data export (acceptable but not mandatory) |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical operation: - Main HFNC unit – 1 - Heated humidifier chamber – 1 - Heated breathing circuits – minimum 5 - Adult nasal cannula – minimum 5 - Pediatric nasal cannula – minimum 5 - Water chambers – minimum 5 - Oxygen tubing – 1 set - Power cable – 1 - Mobile stand or pole mount – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site installation and setup - Functional verification of flow, FiO₂, and temperature accuracy - Alarm system verification - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site clinical training for nurses and respiratory therapists - Training includes operation, cleaning, humidifier maintenance, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel costs - Preventive maintenance included during warranty period - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/