MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.7

| Priority | 1 | Floor | 1F-DR | Quantity | 3EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Incubator (Hybrid + Warmer) | Installation Location | NICU |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The hybrid neonatal incubator with radiant warmer shall provide both enclosed incubator mode and open radiant warmer mode for neonatal intensive care and resuscitation. - The system shall provide precise thermal regulation, humidity control, and safe access to the infant while maintaining a stable microenvironment. - The equipment shall be suitable for premature and critically ill neonates requiring continuous life-supportive thermal care. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Combined incubator and radiant warmer in a single integrated platform - Seamless conversion between closed incubator mode and open warmer mode - Microprocessor-controlled system - Continuous-duty NICU operation - Medical-grade disinfectant-resistant surfaces - Smooth edges and easy cleaning design - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Temperature Control Performance – Incubator Mode [MUST] |  |  |  |  |  |
| - Air temperature control range: 25–38°C or wider - Temperature accuracy: ±0.2°C or better - Temperature uniformity within chamber: ±0.5°C or better - Servo-controlled skin temperature mode with probe - Air temperature mode selectable - Warm-up time to 36°C ≤ 15 minutes - Automatic compensation for door opening or heat loss - Stable temperature maintenance during prolonged operation |  |  |  |  |  |
| - Radiant Warmer Performance – Open Mode [MUST] |  |  |  |  |  |
| - Overhead radiant heating element - Skin servo control with automatic power modulation - Fast response heating - Even heat distribution across mattress surface - Manual power adjustment option - Safe overheating protection |  |  |  |  |  |
| - Humidity Control [MUST] |  |  |  |  |  |
| - Relative humidity control range: minimum 0–90% - Humidity accuracy ±5% or better - Stable humidity maintenance - Large water reservoir ≥ 1 liter - Easy refill without opening hood - Condensation control design |  |  |  |  |  |
| - Air Circulation & Environmental Stability [MUST] |  |  |  |  |  |
| - Forced air circulation system for uniform heat distribution - Airflow designed to minimize drafts - Internal air filtration (bacterial/HEPA equivalent) - Noise level ≤ 45 dB preferred (≤ 50 dB maximum) - Low vibration operation |  |  |  |  |  |
| - Infant Bed & Mattress System [MUST] |  |  |  |  |  |
| - Radiolucent mattress platform for X-ray without moving patient - 360 degree rotatable infant bed - Adjustable tilt (Trendelenburg/Reverse) minimum ±10° - Height of the whole unit should be adjustable - Mattress size appropriate for neonatal care (minimum 65 × 35 cm approx.) - Removable and washable mattress cover - Must have Integrated weighing scale |  |  |  |  |  |
| - Hood & Access [MUST] |  |  |  |  |  |
| - Transparent double-wall hood - Multiple access ports with soft sleeves - Large side doors for emergency procedures - One-hand quick-open mechanism - Secure locking system - Easy removal for cleaning |  |  |  |  |  |
| - Monitoring & Display [MUST] |  |  |  |  |  |
| - 8- inch LCD / LED Color digital touch display showing: - Air temperature - Skin temperature - Humidity - Heater output - Alarm status - Real-time numeric and graphical display - Easy-to-use control interface |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - High/low air temperature alarms - High/low skin temperature alarms - Over-temperature protection - Sensor failure alarm - Power failure alarm - Door open alarm - Water level alarm - Heater malfunction alarm - Audible and visual alarm indicators |  |  |  |  |  |
| - Sensors & Control [MUST] |  |  |  |  |  |
| - Minimum two (2) skin temperature probes included - Redundant temperature sensing - Automatic self-test at startup - Accurate closed-loop servo control |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz - Internal battery backup ≥ 30 minutes for alarm and control continuity - Battery level indicator |  |  |  |  |  |
| - Continuous Operation [MUST] |  |  |  |  |  |
| - Designed for 24-hour continuous NICU use - Stable performance without overheating or drift |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Hybrid incubator/warmer main unit – 1 - Integrated trolley with lockable wheels – 1 - Mattress – 2 - Skin temperature probes – minimum 2 - Humidity water chamber – 1 - IV pole – 1 - Oxygen inlet port – 1 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site installation and calibration - Verification of temperature, humidity, and alarm performance - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for NICU nurses and biomedical engineers - Includes operation, cleaning, humidity maintenance, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel - Preventive maintenance included during warranty - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-19 (Infant Incubators) - CE MDR Regulation (EU) 2017/745 and FDA 510(k) certification - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/