MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.8

| Priority | 1 | Floor | 1F-DR | Quantity | 4EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Infant Radiant Warmer | Installation Location | NICU |  |
| Specification Level |  | High | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The radiant warmer is a thermoregulation system designed to provide a controlled thermal environment for newborns and infants. - Its primary purpose is to stabilize and maintain the patient’s core body temperature at approximately 37 °C through a microprocessor-controlled infrared heating element. - The system is intended for use in delivery rooms, neonatal intensive care units (NICUs), and pediatric wards to facilitate routine care, resuscitation, and minor surgical procedures while preventing neonatal cold stress. - It shall support both "Servo" (skin-temperature-led) and "Manual" (user-defined output) modes to ensure clinical flexibility in both stable and emergency obstetric or neonatal scenarios. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - Control System: Microprocessor-controlled radiant warmer for continuous optimal thermal support with an advanced safety alarm system. - Operational Modes: Must feature pre-warm, manual, and baby servo operation modes. - Design: Complete unit with an attached baby bassinet and all standard accessories. - Mobility: Mobile trolley-based radiant warmer with lockable castors. - Construction: Must have a double handle for easy transportation and foldable, transparent, high-quality side panels for easy access to the newborn. |  |  |  |  |  |
| - Technical Specifications & Performance [MUST] |  |  |  |  |  |
| - Heating Element: Ceramic or quartz heaters with heating element output ≥ 600 W or equivalent thermal performance, designed to deliver uniform heat distribution throughout the newborn’s body. - Heater Module: Swivel/rotation capability to allow for X-ray procedures with minimal stress to the infant. - Temperature Range (Baby Mode): Adjustable from approximately 32°C to 39°C with a sensor accuracy of ± 1°C or better. - Manual Mode: Heater output can be converted from servo mode to manual mode for non-stop working. - Imaging: Must feature an X-ray cassette tray accessible from the front side of the bed. |  |  |  |  |  |
| - Infant Bed & Mattress System [MUST] |  |  |  |  |  |
| - Must support a newborn weight capacity ≥ 10 kg, with a removable, waterproof, and washable mattress. |  |  |  |  |  |
| - Monitoring & Display [MUST] |  |  |  |  |  |
| - Large and bright LED/LCD display showing: - Air temperature - Skin temperature - Heater output - Alarm status - Real-time numeric and graphical display - Easy-to-use control interface - Display can be converted from °C to °F |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - High/low air temperature alarms - High/low skin temperature alarms - Sensor failure alarm - Power failure alarm - Audible and visual alarm indicators |  |  |  |  |  |
| - Sensors & Control [MUST] |  |  |  |  |  |
| - Minimum two (2) skin temperature probes included - Accurate closed-loop servo control |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC 100–240 V, 50/60 Hz with appropriate power cable |  |  |  |  |  |
| - Continuous Operation [MUST] |  |  |  |  |  |
| - Designed for 24-hour continuous NICU use - Stable performance without overheating or drift |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include: - Infant Radiant warmer main unit – 1 - Integrated trolley with lockable wheels – 1 - Mattress – 2 - Skin temperature probes – minimum 2 - IV pole – 1 - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - On-site installation and calibration - Verification of temperature, humidity, and alarm performance - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for NICU nurses and biomedical engineers - Includes operation, cleaning, maintenance, and troubleshooting |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years after acceptance - Covers parts, labor, and travel - Preventive maintenance included during warranty - Response time ≤ 48 hours - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/END/