MEDICAL EQUIPMENT DETAILED SPECIFICATION

Specification No.9

| Priority | 1 | Floor | 1F-DR | Quantity | 2EA |
| --- | --- | --- | --- | --- | --- |
| Equipment Name |  | Phototherapy System (LED) | Installation Location | NICU |  |
| Specification Level |  | Mid | Remarks |  |  |
| General Specification |  | The following items are basic requirements; equivalent or higher specifications with similar functionality are also acceptable. |  |  |  |
| - Scope and Intended Use |  |  |  |  |  |
| - The LED phototherapy system shall be used for treatment of neonatal hyperbilirubinemia (jaundice) by providing effective blue light irradiation. - The system shall deliver uniform, stable, and safe light intensity suitable for routine neonatal therapy in delivery rooms and newborn care areas. - Equivalent or higher performance systems are acceptable. Vendor-neutral and performance-based specifications shall apply. |  |  |  |  |  |
| - General Requirements [MUST] |  |  |  |  |  |
| - LED-based phototherapy system - Overhead or mobile stand type suitable for neonatal beds/incubators - Quiet operation without active cooling noise or excessive heat - Easy-to-clean and disinfectant-resistant surfaces - Stable mechanical structure with smooth height adjustment - The equipment shall be brand new, unused, and currently in production at the time of bid submission. Refurbished or discontinued models shall not be accepted. |  |  |  |  |  |
| - Irradiance Performance [MUST] |  |  |  |  |  |
| - Wavelength range: approximately 430–490 nm (blue light spectrum suitable for bilirubin reduction) - Peak wavelength around 460–470 nm - Irradiance intensity: ≥30 µW/cm²/nm (high-intensity mode acceptable) at mattress level - Uniform irradiance distribution across treatment area (variation ≤ ±20%) - Stable output without significant decay during operation Adjustable light intensity levels |  |  |  |  |  |
| - Treatment Area Coverage [MUST] |  |  |  |  |  |
| - Coverage area suitable for full-body neonatal treatment - Light field adequate for newborn positioned in bassinet/incubator - Adjustable height or angle for proper positioning |  |  |  |  |  |
| - Thermal & Noise Performance [MUST] |  |  |  |  |  |
| - Low heat emission to avoid neonatal overheating - Surface temperature safe for neonatal use - Quiet operation suitable for NICU environment (Noise level ≤ 50–55 dB) |  |  |  |  |  |
| - Control & Interface [MUST] |  |  |  |  |  |
| - Digital display of operating time - Irradiance level selection (low/high modes or adjustable) - Treatment timer - Hour meter for LED usage - Easy-to-use controls |  |  |  |  |  |
| - Safety Features [MUST] |  |  |  |  |  |
| - Over-temperature protection - Stable mechanical locking of arm/stand - Electrical leakage protection compliant with medical safety standards - Light shield or eye protection support for patient use |  |  |  |  |  |
| - Power Requirements [MUST] |  |  |  |  |  |
| - AC power 100–240 V, 50/60 Hz - Low power consumption LED technology - Stable operation during continuous long-term therapy |  |  |  |  |  |
| - LED Life & Maintenance [MUST] |  |  |  |  |  |
| - LED life ≥ 20,000–30,000 hours - Low maintenance design - Radiometer or irradiance verification method recommended - Easy cleaning/disinfection |  |  |  |  |  |
| - Accessories & Configuration [MUST] |  |  |  |  |  |
| - Each unit shall include all items required for immediate clinical use: - Phototherapy light head – 1 - Mobile stand or mounting arm – 1 - Height adjustment mechanism – integrated - Protective light shield/cover – 1 - Eye protection goggles for neonate – minimum 5 sets - Power cable – 1 - User manual and service manual – 1 set |  |  |  |  |  |
| - Optional Items (price separately) |  |  |  |  |  |
| - Additional eye protection sets - Radiometer for irradiance measurement - Spare LED modules - Battery backup module (if available) |  |  |  |  |  |
| - Installation & Commissioning [MUST] |  |  |  |  |  |
| - Delivery and on-site setup - Functional testing of irradiance and timer - Electrical safety verification - Acceptance testing checklist completed |  |  |  |  |  |
| - Training [MUST] |  |  |  |  |  |
| - On-site training for nurses and neonatal staff - Training includes correct positioning, safety precautions, and cleaning procedures |  |  |  |  |  |
| - Warranty & After-Sales Service [MUST] |  |  |  |  |  |
| - Comprehensive warranty minimum 3 years - Covers parts, labor, and travel costs - Response time ≤ 48 hours - Preventive maintenance included during warranty - Spare parts availability ≥ 7 years |  |  |  |  |  |
| - Compliance & Certification [MUST] |  |  |  |  |  |
| - IEC 60601-1 (Medical electrical equipment – General requirements for basic safety and essential performance) - IEC 60601-1-2 (Electromagnetic compatibility requirements) - IEC 60601-2-50 (Phototherapy equipment) - CE marking or FDA clearance or equivalent international regulatory approval - Manufacturer shall be ISO 13485 certified. |  |  |  |  |  |

/End/